## Slide 1
Q1: How to calibrate the double side SSD in our experiment?
Q2: 电荷重心法

## Slide 3
1. Basic Knowledges of ∆E1- ∆E2-E Detector

## Slide 4
Date: 2018-10-23
Source: https://groups.nscl.msu.edu/hira/05049/LASSA_Calibrations.pdf
Calibration on Silicon Detectors
● Th228, characterized by eight well-known alpha energies, and alpha peak are separated by at least 0.1MeV
● If a dE silicon detector is placed in front of the double sided E silicon detector, the dE detector can be calibrated by 228Th source as discussed above; while pin source is used to calibrate E detector

## Slide 5
Date: 2018-10-27
Source: Talk with Daniele
Si-Si-CsI detector data analysis procedures：
1. Energy calibrations of Si detector with alpha source
2. Charge sharing method to determine the position of the hits: if the hit happens between two strips
3. Calibrations of CsI
4. Pixelization
5. Multiplicity analysis
6. Double hit or multi-hit
7. Particle Identification

## Slide 6
Date: 2018-10-30
Source: Talk with Daniele
Double-sided Silicon Strip Detector
1. What is a “pixel” ?
● Double-sided silicon strip detector can read out signal on both the
sides. The back strips are perpendicular to the front strips. A “pixel”
means the overlap of a font strip and a back strip.
● Double-sided SSD also called “matrix detector”
2. What does a pixel works as a position sensitive detector?
● The position of SSD is determined by the pixel. When charge particle pass
through a SSD, ideally, there will be a same a energy signal on both a front strip
and a back strip, which determines the position of the particle hits the detector.
● Charge sharing effect: means the charge generated by a particle is collected
by two or more than two strips. In this case, we should find the strip with the
largest energy, this is the strip that the particle hitting or the strip nearest to the
hitting point. Then, sum up the energy in the adjacent strips to this strip.
● 对于 Charge sharing，需要采用电荷重心法来处理。

## Slide 7
3. What is multiple hit?
(1) Multiple hit means there are multiple particles hits the SSD with one readout period (no need hit
simultaneously). Silicon strip can not assign the hit positions clearly.
(2) 如何筛选出“多击事件”事件呢？
● 在pixelization过程中：当前面某个strip探测到一个能量信号时，原则上需要用这个strip跟
背面的32个strip进行组合，如果背面只有一个strip(考虑charge sharing之后)也同时探测到
信号，并且两个能量信号相同(或接近相同)，那么这两个strip对应的“pixel”就是入射粒
子的位置. 这种情况下，前面和背面总能match上.
● 如果后面不止有一个strip有信号，这种情况下，背面的每个信号会比前面小得多。前面和
背面match不上！ 对于这种“多击事件”，通常我们通常直接扔掉！！

## Slide 11
2. Learn from Papers

## Slide 12
Date: 2018-10-28
Source: https://doi.org/10.1016/j.nima.2017.12.082

## Slide 13
Date: 2018-10-28
Source: https://doi.org/10.1016/j.nima.2017.12.082
1. Introduction
(1) The dead layer on the surface protects the active silicon wafer of the detector.
(2) HiRA telescope: the High Resolution Array
configuration:
● a 65- single-sided silicon strip detector (“”, manufactured by Micron Semiconductor)
● a 1500- double sided silicon strip detector (“E”, manufactured by Micron Semiconductor)
● an array of 4 CsI scintillator crystals, read-out by silicon photodiodes
(3) At the front of each telescope: a thin mylar foil pretecting the sensitive silicon detectors, and also completing a Faraday cage around the detectors to minimize the noise.
a Faraday Cage is just an electronic isolation chamber. Or even more simply it’s a special container that prevents electrical signals or waves from passing through it.

## Slide 14
Date: 2018-10-28
Source: https://doi.org/10.1016/j.nima.2017.12.082
(4) The strips on the opposite sides of E detector are perpendicular to each other, so detecting a particle in both a front strip and a back strip results in a single 1.95mm “pixel”
(5) What can HiRA do?
● Excellent energy resolution => to detect energy
● Excellent position resolution => to detect moving direction
● particle identification => to detect mass && charge
we can get E(energy) and

## Slide 15
Date: 2018-10-28
Source: https://doi.org/10.1016/j.nima.2017.12.082
2. Dead layer thickness determination
(1) Using Th228 alpha source to calibrate the HiRA E and ∆E detectors. 3 advantages:
● Six clearly separated peaks with energy from about 5 MeV to 9MeV
● Th228 sources of various strengths are commercially available
● Th228 has a relatively long half-life(1.9 year)
(2) How to get the Th228 source?
● Electroplated onto a platinum surface
● Then fixed in an aluminum holder with a gold window.
(3) Ideally, if an alpha source can be inserted between the ∆E and E detectors,
the E detector can be calibrated without disassembling the HiRA telescope.
(4) The pin source is primarily Pb212 (one of the isotopes from the decay chain
of Th228)
(5) The HiRA telescope was designed so that the pin source mounted in the
frame sits 3.2mm above the surface of the E detector when the frame
is inserted into the slot.

## Slide 16
Date: 2018-10-28
Source: https://doi.org/10.1016/j.nima.2017.12.082
(6) The idea of testing the thickness of the dead layer
● The dead layer is the only potential cause of energy loss between the source and active
volume, as alpha particles from the pin do not need to pass through a gold window or
Mylar foil to reach the detector.
● Some assumptions:
<1> Assuming an effective dead layer of pure silicon
<2> Assuming a uniform dead layer, means a constant thickness of the dead layer
<3> Assuming a constant value for energy loss, since the thickness of the dead layer is very small.
● Basic formula used for extraction of T
E: detected energy in E detector
: the initial energy of alpha particle
: emission angle
: stopping power of an alpha particle in Si at energy
(assumed to be a constant)
Observable quantities: E ,

## Slide 17
Date: 2018-10-28
Source: https://doi.org/10.1016/j.nima.2017.12.082
● Only two peaks in the pin source data:
6.062 MeV && 8.785MeV
8.785MeV
Two unresolved peaks: 6.051MeV(25.1%) & 6.090MeV(9.75%)
Combined via a weighted sum:
6.062MeV
● Steps to determine the thickness of dead layer
<1> Initial calibration using Th228 source, with a reasonable guess
of
<2> To find the central pixel: the pixel closest to the pin source.
● The central pixel provides a good approximation to the precise
location of the pin source, which has relatively large solid angle
coverage.
● Since the emission of alpha particle is isotropous, the central
pixel should have more counts
● The central pixel is at the intersection of the central front strip
and central back strip

## Slide 18
Date: 2018-10-28
Source: https://doi.org/10.1016/j.nima.2017.12.082
<3> Fit for each peak of each pixel
: the front strip number of the pixel
: the back strip number of the pixel
: the front strip number of the central pixel (the front pin source position)
: the back strip number of the central pixel (the back pin source position)
: width of each strip(fixed to 1.95mm)
h: the distance from the pin source to the detector surface (fixed to 3.2mm)
: initial energy
Only 2 free parameters: T and

## Slide 19
Date: 2018-10-28
Source: https://doi.org/10.1016/j.nima.2017.12.082
3. Si detector thickness determination
(1) Dead layer has a smaller effect on relatively high energy particles that punch through the entire E
detector. In this case, it’s important to understand the overall thickness of E since CsI are typically
calibrated using particles that punch-through the E.
(2) Two ways to calibrate a CsI crystal:
<1> Kinematics points: elastically scatter light charge particle at known energies into the crystal
● projectile: proton, target: CH2
reaction channel:
● In this experiment, since the kinematic relationship between scattering angle and energy is weak,
this results in one elastically scattered proton calibration per crystal.
● For kinematic points: the higher energy peak corresponds to proton
elastic scattering off carbon, and the lower peak is from inelastic
scattering ().

## Slide 20
Date: 2018-10-28
Source: https://doi.org/10.1016/j.nima.2017.12.082
<2> ∆E-E points (Energy loss method)
● Two important notes:
* This method relies on the assumption that the CsI detector response is linear at low energy (confired by WMU data)
* Requiring that only data from proton hitting the detector be used.
(3) In this energy range, kinematic points and ∆E-E points should be collinear. Since the ∆E-E points are sensitive to the E thickness, the E thickness can be extracted.
(4) Summarize the idea for Si detector thickness determination
<1>Use kinematic points to constrain the calibration at high energy
<2> ∆E-E points should be linear at low energy and be collinear with
kinematic points
<3> Adjust the E thickness to satisfy the condition in <2>

## Slide 21
Date:
Source:

## Slide 22
Date:
Source:

## Slide 26
3. Scintillator—inorganic crystal

## Slide 27
A few recalls: Mechanism of light production in CsI(Tl).
(alkali halide inorganic crystal)
Valence band (here the electrons are bound to lattice sites)
Conduction band (here electrons are free to move along the crystal)
Band gap (~10eV)
photon
Pure crystal
Valence band (here the electrons are bound to lattice sites)
Conduction band (here electrons are free to move along the crystal)
Band gap (~10eV)
Doped crystal
(impurities are added, Tl or Na)
Forbidden band (an electron cannot be found here)
energy absorption
Accessible states are created in the forbidden band
(activator)
energy absorption
optical photon
After ionizing, a free electron is located in the conduction band and a hole is located in the valence band. The electron can return to the valence band emitting a photon (normally not in the visible range).
When the impurity is present, the hole will drift to an activation center and ionize it (because the ionization energy is much less for the activator centers than for a lattice site). When the electron, free to move in the conduction band, encounters the ionized activator, a recombination process takes place. A neutral activator is created in one of its excited states  decay via the emission of an optical photon.
hole
electron
hole
electron
hole drift
ionized activator (luminescence center)
electron recombination
Date: 2018-10-31
Source: Daniele’s group meeting talk
Mechanism of light production in CsI(Tl) and light response

## Slide 28
The interaction of heavy charged particles with the CsI(Tl) occurs mainly via ionization mechanisms (and, in minor quantity, via interactions with nuclei in the lattice).
primary ionization column (high-density of charge carriers)
Incident ionizing particle
e-
e-
e-
e-
e-
The efficiency of light production within the ionization column is limited by quenching effects:
Ideal light response
Quenched light response

## Slide 29
Furthermore, the light response is not a constant with the particle’s energy. A unified treatment of light response was proposed by Birks for alpha-particles in organic scintillators (J.B. Birks, Scintillation Counters, Pergamon Press, Oxford 1960, p. 93) and is normally applied also in the case of inorganic scintillators.
electronic stopping power
Scintillation efficiency factor
Quenching factor
This formula can be analytically integrated by considering the following approximation valid for low energy:

## Slide 30
This leads to the following formulations of the CsI(Tl) light response:
D. Horn et al., Nucl. Instr. Meth A 320 (1992) 273
N. Colonna et al., Nucl. Instr. Meth. A 321 (1992) 529
P.F. Mastinu et al., Nucl. Instr. Meth A 338 (1994) 419
Y. Larochelle et al., Nucl. Instr. Meth. A 348 (1994) 167

## Slide 31
However, the scintillation efficiency dL/dE is not a function of the electronic stopping power only, but rather of a composed set of discrete functions: A. Meyer and R.B. Murray, Phys. Rev. 128 (1962) 98.
The figure, from M. Parlog et al., Nucl. Instr. Meth. A 482 (2002) 674, shows that for the same average stopping power heavier isotopes have a greater light efficiency.

## Slide 32
This is the effect of the so-called d-electrons (or knock-on electrons), i.e. modestly energetic electrons (~1keV) that are scattered outside of the primary ionization column caused by the incident ionizing radiation.
primary ionization column (high-density of charge carriers)
Incident ionizing particle
e-
e-
e-
e-
e-
e-
e-
e-
“virgin” region of the crystal
“virgin” region of the crystal
Knock-on electrons have access to a “virgin” region of the crystal, outside of the highly ionized region, where quenching effects don’t occur. This is a high-efficiency region of the crystal. The Birks formula has to be changed as follows:
Fractional energy loss. This is the fraction of energy carried off by d-rays. Non-zero only for

## Slide 33
The fractional energy loss is non-zero only when the energy exceed the threshold of delta-ray emission and it is more important for heavier isotopes (since light isotopes have a small efficiency of delta-ray production) and at higher energies. An empirical model for fractional energy loss calculation and its dependency on E/A was first given by Meyer and Murray.
The most recent formulation is by Parlog (Recombination and nuclear quenching model). This leads to an approximate formula:
Fractional energy loss (this is put artificially to 0 for )
Threshold of delta-ray production
Example from O. Lopez et al., Nucl. Instr. Meth. A 884 (2018) 140

## Slide 34
Quality of fit results from the literature
M. Parlog et al., Nucl. Instr. Meth. A 482 (2002) 693
Applied down to Z=5.

## Slide 35
Quality of fit results from the literature
D. Horn et al., Nucl. Instr. Meth A 320 (1992) 273
General formula valid over a range of light ions.
Limited energy range available for protons.
Y. Larochelle et al., Nucl. Instr. Meth. A 348 (1994) 167

## Slide 36
Quality of fit results from the literature
A. Wagner et al., Nucl. Instr. Meth. A 456 (2001) 290
A better fit quality is obtained by fitting isotopes of each Z with an individual empirical formula.
For Z=3 
For Z=2 
For Z=1 
Only 1 calibration point is present for 1H!

## Slide 37
Quality of fit results from the literature
C.J.W. Twenhofel et al., Nucl. Instr. Meth B 51 (1990) 58
Dedicated apparatus to measure light response of light particle exclusively.
Saturation effects observed for Z<=3, more important for Z=1. The onset of this effect is between 20 and 30 keVcm2/mg for Z=1, between 30 and 40 keVcm2/mg for Z=2 and about 70 keVcm2/mg for Z=3.
Saturation effect
Separate fits for Z<3 because of saturation effects.

## Slide 38
Quality of fit results from the literature
G. Viesti et al., Nucl. Instr. Meth A 252 (1986) 75
Linear fit on Z=1 based on few calibration points. The same data was interpreted in terms of the previously explained saturation effects by Twenhofel et al.

## Slide 39
Quality of fit results from the literature
S. Aiello et al., Nucl. Instr. Meth. A 369 (1996) 50-54
Similar saturation effect possibly observed also by Aiello et al.
Separate fit for Z=1 because of saturation effects.

## Slide 40
It is clear that very few and fragmentary data is reported in the literature for Z=1 light response in CsI(Tl) (and inorganic scintillators in general).
Since we are interested in proton energy spectra over a wide range of energies (and for the first time we have good quality data of protons in a wide energy range), it is mandatory to explore the possible existence of saturation effects observed in the literature and build a proper fit formula.
L (V)
E (MeV)
p,d,t
3He
4He
6,7Li
7Be