## Slide 1
Tutorials for HiRA CsICalibrations

## Slide 2
DE – E points
Ca+CH2 Data
Punch Through points
Kinematics Line
CsILength
Cuts, Projection
Kinematics points
CsI Calibrations
DE-E Data
Fragmentation Data
Step 1
Step 2
Step 3
Step 4
Step 5
WMU Data
Si Calibration
Beam Geometry
affect
affect
What we need to do CsI Calibrations?

## Slide 3
I. How to get different kinds of datasets?
1. How to get ∆E-E points?
2. How to get kinematic points?
3. How to get punch-through points?

## Slide 4
1. How to get ∆E-E points?
Two methods:
1. Projection method
2. DEEFIT method
Notes: ∆E-E points depend on the calibrations of silicon strip!

## Slide 5
1. Projection method
Notes: For those particles travel through silicon strips and stop in the CsI, from the ∆E-E spectra, we know the energy loss[∆E(MeV), calibrated] in silicon strips, and the energy deposited in CsI [ECsI(ch)].
Step 1: (macro to be used: 00DEEProfilingPoints/ macro_PunchThroughCuts.cxx)
(1) Do a cut for each particle in the ∆E-E spectra
Step 2: (macro to be used: 00DEEProfilingPoints/ macro_CsIDEEProfiling.cpp)
(1) Within the cut region, take samples points of a constant bin step in X-axis, make projection to Y-axis, do Gaussian fit.
=> ∆E(MeV), err ∆E(MeV) && ECsI(ch), errECsI=0.5(ch) (set by hand)
(2) Use pulser calibration data to convert energy deposited in CsI from channels [ ECsI(ch)] to voltages [CsIV(V)]
[Notes: pulser calibration data can be attained by using precision pulse generator]
=> ∆E(MeV), err∆E(MeV) && ECsI(ch), errECsI=0.5(ch) && CsIV, errCsIV
Step 3: (macro to be used: 00DEEProfilingPoints/ macro_CsIPIDCalibrationPointsFromProfiling.cpp)
Use LISE module to calculate the energy deposited in CsI in MeV [CsIE(MeV)]
=> CsIV, errCsIV && CsIE, errCsIE

## Slide 6
Step 4: To estimate the the errors of ∆E-E points
For a certain channel number of CsIch, the error err_CsIch=0.5, set by hand
(2) For errors of voltage:
CsIV = CsICalibrationModule.GetVoltageValue(CsIch, telnum, csinum);
err_CsIV = CsICalibrationModule.GetVoltageValue(CsIch+err_CsIch, telnum, csinum) - CsIV ;
(3) For Gaussian fit
ESi = fit_result->GetParameter(1);
// the center of the Gaussian peak
err_ESi = fit_result->GetParError(1);
// the error of the center of Gaussian fit
(4) For errors of CsIE
∆E(MeV)
ECsI(Ch)
CsIch
ESi(∆E)

## Slide 7
Notice 1:
∆E(MeV)
ECsI(Ch)
CsIch
Esi(∆E)
MidEnergy
fewer points
LowEnergy
more points
1. Since low energy points (smaller ECsI) are more reliable, I would like to take a little bit more points at low energy regoin

## Slide 8
Notice 2:
1. The way to calculate the energy deposited in CsI by using the energy loss in silicon (∆E)
2. An example of the calibration of CsI

## Slide 9
2. DEEFIT method
DEEFIT: A tool for charge and mass identification in Si-CsI matrices od the CHIMERA detector
See the readme “docdeefit.1.1.pdf” inside the DEEFIT folder
The procedures are based on the followings:
(1) Obtain a set of raw data for a set of telescopes with a well defined ∆E-E spectra
(2) Draw a set of sampling clicks for selected Z and M in the Si-CsI matrix
(3) Fit the data(click points) with a multi-dimensional function
(4) Get charge and mass identification by an iterative procedure based on the previous fit result.
(5) Check the goodness of the results on charge and mass spectra.
How to run DEEFIT:
For Standalone version (I use this version):
simply sun “deefit_S”: ./deefit_S

## Slide 10
Step 1: Open file
Click 1
Click 2
Click 3
Step 2: Plot
Click 1
Step 1: Chose the correct data to be fitted
Step2: Chose a certain telescope to plot

## Slide 11
Step 3: Set charge Z and mass A, draw the clicks for each particle
Step3: chose a certain particle to draw
Click 1
Click the center of the line
Then, the clicks will be saved in the list.
Click 2

## Slide 12
Step 4: After drawing the clicks for all particles, then fit
Click 1
Two fit function, the one with14 parameters is recommended
Click 2
Click 3 Init Fit
Click 4
Click Fit, until the ChiSquare is minimum!

## Slide 13
Step 5: Check the fit result
Chose a certain particle
Click 1
We can see if the fit and the clicks match or not

## Slide 14
Step 6: One of the most powerful function of DEEFIT-------- “particle prediction ”
According to the fit result, triton should be here!
In this tutorial, we just draw the clicks for proton, according to the fit result, we can predict where are the other particles. Of cause, with more particles to do the fit, we will get more accurate predictions!

## Slide 15
Step 7: Delete one click and redo the fit
If we don’t satisfied with the fit result, then delete some clicks, and redo the clicks

## Slide 16
Step 8: Analysis
Mass distributions
Sea “docdeefit1.1.pdf” for detials

## Slide 17
Step 9: If the fit result is good enough(the ChiSqurae is small enough ), then save fit parameters
Click 1
We can also load the parameters next time when we want to improve the fit!
Step 10: Save the clicks
Click 1
We can also load the click file next time when we want to improve the fit!

## Slide 18
Use the DEEFIT parameters to get ∆E-E points:

## Slide 19
To estimate the errors of ∆E-E points:
(1) For errors of voltage:
CsIV = CsICalibrationModule.GetVoltageValue(XChannel, numtel, numcsi);
err_CsIV = errPercentageCsIV * CsIV; // add 1% errors by hand
(2) For energy loss in Si:
ESi = func14(input, par);
err_ESi = errPercentageESi *ESi; // add 1% errors by hand
(3) For errors of CsIE

## Slide 21
2. How to get kinematic points?

## Slide 22
Step 1: Build Histograms for Kinematics Lines
The raw kinematics data can be saved as a tree with different branches, e.g. HiRA.fmulti, HiRA.fnumtel, HiRA.fnumcsi, HiRA.fEnergycsi, HiRA.fTheta …
We need HiRA.fEnergycsi && HiRA.fTheta to build the histograms of kinematics lines
HiRA.fTheta(Degree)
HiRA.fEnergycsi (ch)
kinematics lines

## Slide 23
Extension: what is kinematics line?

## Slide 31
3. How to get punch-through points?

## Slide 37
II. How to fit the datasets simutaneosly?

## Slide 43
Some basic questions?

## Slide 44
Q1: How to build the ∆E-E spectra?

## Slide 45
Q2: How to build kinematics data ?